## Slide 1: Study and development of a Validation Facility for the SST camera of the CTA observatory
Candidate: Isabella Sofia
Supervisor: Prof. Andrea Chiavassa
Co-supervisor: Dr. Davide Depaoli
Co-examiner: Prof. A. Ferretti

## Slide 2: Summary
Brief γ-ray astronomy and IACT introduction
  CTA and SST, CHEC-S
The validation facility at MPIK
  Goals
  Experimental set-up
  Study of components
    The Night Sky Background (NSB)
    The Filter Wheel and its calibration
    The laser beam model (Robot Arm)
CHEC-S: an SST prototype
  CHEC-S in the DarkBox
    Trigger Delay
    HV settings
Conclusions and outlooks

## Slide 3: Neutral cosmic messengers: gamma rays
Neutral messengers: photons
  Massless
  Not influenced by magnetic fields
  Point back to their sources
Scientific goals
  Identify cosmic rays’ sources
  Study and understand the origin of such cosmic rays
  Probe peculiar environments (SNRs, pulsars, AGNs, PWNs, …)
  New physics (e.g. Dark Matter)
Methods of detection
  Direct: satellite experiments
  Indirect: ground-based
Crab Nebula Pulsar (top) and Tycho’s SNR (bottom)

## Slide 4: IACT technology
Imaging Atmospheric Cherenkov Telescopes
Ground-based experiments
Use the Earth’s atmosphere as a «sampling calorimeter»
Detect the Cherenkov light produced by electromagnetic showers
It is possible to reconstruct:
  Direction
  Energy
Allows excellent γ-hadron separation
Low Duty Cycle
Limited field of view

## Slide 5: The CTA Observatory
Cherenkov Telescope Array
Next generation IACT observatory
A northern and a southern hemisphere site to cover the whole sky
  Roque de los Muchachos, la Palma, Canary Islands, Spain
  Chile
Wide energy range: 20 GeV - 300 TeV
Three telescope classes with different mirror diameter
  Small-sized (SST) – 4 m
  Medium-sized (MST) – 12 m
  Large-sized (LST) – 23 m

## Slide 6: SST – Small-Sized Telescopes
On southern site only
Sensitivity range from ≈ 1 TeV to 300 TeV
Three different SST structures were proposed (SST1-M, ASTRI-Horn, CHEC)
Final design: ASTRI (telescope structure) + CHEC-S (detector)
Dual-mirror Schwarzschild–Couder design (only SSTs)
SST camera working rate of ≈ 600 Hz
ASTRI telescope structure
CHEC-S

## Slide 7: CHEC-S
Compact High Energy Camera
Proposed prototype for the dual-mirror SST design
SiPM-based (Hamamatsu S12642-1616PA-50 SiPMs)
2048 pixels distributed on 32 modules, each comprising 64 pixels of 6 mm ×6 mm
Final SST Camera design will be CHEC-S-like, plus…
  Improved cooling and air flow
  Glass window
  Filter on window
  New SiPM technology (S14521-1720)
J. Zorn
New flat window
CHEC-S prototype

## Slide 8: Validation Facility – «DarkBox»
≈ 42 SST Cameras to be developed
Validation for each newly produced camera is needed
The «Validation Facility» (DarkBox) at MPIK will
  Provide a dark room for measurements
  Simulate a certain range of background light coming from the sky (NSB)
  Simulate the Cherenkov light (Laser)
  Be completely remotely-controlled by the user (via the DarkBox Manager Software)
  Be a stable and well-known environment to test each SST camera
Validation Facility at MPIK Laboratories
Optical set-up

## Slide 9: Simulating the NSB – set-up
Use of a DC LED
NSB is a diffused radiation  use of a 20° diffuser
Use of two different ND filters in order to cover a range between ≈10 MHz to ≈ 20 GHz
Powermeter at NSB for feedback loop
Photon rate on SiPM VS Power at NSB source

## Slide 10: Simulating the Cherenkov light - laser
Use of a 400 nm UV Picoquant laser
Use of two diffusers (20 + 5 °)
Required uniformity of the beam of 2 %
Center of beam should be reasonably at the center of the camera
Goal: simulate over a continuous range of photons arriving at detector (e.g. )

## Slide 11: Providing a wide range of photon rate – the Filter Wheel
Laser light intensity should remain constant (stability issues)
Filter Wheel: continuously variable neutral density filters that provide attenuation
  Controlled with the DBM Software (via ethernet)
  Attenuation can be adjusted linearly via rotation
  Attenuation covers 4 orders of magnitude (OD 0.04 – 4.0)
  Calibrating the filter wheel means knowing how many photons we can expect at camera position at each step on the FW at the SST working rate
  To obtain this: study the filter wheel range, i.e. “calibrate” the filter wheel

## Slide 12: Filter Wheel «calibration» - procedure
Powermeter measurements at camera position (3 m) at 600 Hz are quite difficult (powermeter sensitivity ≈ pW)
Hence, two sets of data are needed
  One at the working rate (600 Hz)
  Another one at a higher rate (1 MHz)
With these, we can obtain the expected number of photons at camera position at working rate for each FW position
Interpolate the result on all FW positions that yield at least 1 photon
Save the interpolated values in a calibration file to be included in the DarkBox Manager Software.
Power measured at laser position

## Slide 13: FW calibration procedure
Number of photons at 3 m from source, 600 Hz
Transmission (laser measures)
Number of photons at 3 m from source, 1 MHz
Dynamic range of laser source: 1 γ → 4000 γ

## Slide 14: Procedure automatization
The facility will often be monitored: need to automatize this procedure in order to have always the updated configurations
Implementation of «Calibrate Filter Wheel» in the DarkBoxManager
The user can, at anytime, run the procedure directly from the DarkBoxManager GUI. Said procedure will
  Take the needed runs of data
  Perform the described analysis
  Return a new calibration file to be set

## Slide 15: Beam Model
Uniformity requirement at 3 m from source
Use a Robot Arm to scan the laser beam on a circular area of maximum radius of 0.30 m (step 1 cm)
Tried different optical combinations to obtain uniformity at 3 m
Final combination: 20° + 5° diffusers and beam expander

## Slide 16: 2D scans
1 m
2 m
3 m
With final optical combination uniformity at 3 m is about 1 %
We can study the resulting beam
  Is it reasonably centred?
  How do the optical diffusers work?
  Are we able to describe the laser beam in some way near the source?
 several fits, both one and two dimensional
   Which function works best?

## Slide 17: Starting point – data
Slicing along y = 0
Slicing along x = 0
Idea
  The laser light is emitted as a gaussian beam
  Diffusers are «circular tophat» diffusers

## Slide 18: Supergaussian distribution
The «Super-gaussian function» has a flat top and a gaussian fall-off
How does it work on data?
Data: slice along y = 0
A supergaussian distribution seems to work quite well on our experimental data

## Slide 19: Results @ 1 m – 2D

## Slide 20: CHEC-S in the DarkBox
Experimental set-up
  Camera positioned on a rotational stage
  Indipendent power supply
  External chiller
  Data carried by an optical fiber
  Always used external triggering
Goals
  «Take back to life» the camera
  Find the signals (i.e. trigger delay)
  Update HV settings (last config file: commissioning in Sicily, 2019)
  In a nutshell: put to test CHEC-S and the DarkBox set-up

## Slide 21: CHEC-S - Delay
Use of various cables to externally trigger the camera  need to find the right «trigger delay» to be set for next measurements, i.e. find the peak of the signal
Implementation of a «delay scan»
Quick analysis to find the peak and visualize it, potentially pixel per pixel
Delay of 300 ns
Python script can be run at anytime if set-up is modified
Here also some «dead» pixels are expected…

## Slide 22: CHEC-S – HV settings
New HV settings
(input of 2000 photons)
Uniformity 10.17%
Amplitude matching: equalize the response to a Cherenkov or laser signal of same intensity between all pixels.
Iterative procedure
Should also improve the extracted charge uniformity between pixels of camera
Before HV settings update
(input of 2000 photons)
Uniformity 17.52 %

## Slide 23: Conclusions and outlooks
Pretty good knowledge of the Validation Facility and its components
  Dynamic range of laser source (FW calibration)
  Description of laser beam
  Improved knowledge of set-up
Fully working experimental set-up to host new SST cameras and/or other detectors
  Fully working Managing Software for next users
Operating CHEC-S Camera
  Known set-up and experimental settings (trigger delay, HV configuration) + scripts obtain them at anytime
Next steps… Use the DarkBox!
  New 64-pixels module for the final SST Camera
Some examples…
  gain calibration
  pulse shape and timing
  gain matching
  PDE measurement of all 64 SiPMs

## Slide 24: Bibliography and Sitography
«Development of a SiPM-based camera for the Large Size Telescope of CTA and study of its performance», D. Depaoli
«Cherenkov Camera and Analysis Development for Highest-Energy Gamma-Ray Astronomy», J. Zorn
«Commissioning and Performance of CHEC-S – a compact high-energy camera for the Cherenkov Telescope Array», J. Watson, J. Zorn
“The Small-Sized Telescopes for the Southern Site of the Cherenkov Telescope Array», R. White
https://www.rp-photonics.com/flat_top_beams.html

## Slide 25
Thanks for the attention

## Slide 26: Back-up

## Slide 27: Cosmic Rays and showers
Hillas plot
Cherenkov light as e.m. shower develops

## Slide 28: γ-rays production
Understanding the γ-rays production
Leptonic models
  Self-Synchrotron Compton (S.S.C.)
    Synchrotron emission from accelerated electrons
    Inverse Compton
Hadronic models
  Photoproduction
    
Spectral Energy Distribution

## Slide 29: Reconstructing the shower
Analysis of VHE γ-ray observations: from raw data (measured charges in individual camera pixels) via camera images (characterised by e.g. the Hillas length, width, and amplitude) to final event lists, spectra, and images of sources
  Image cleaning (removing pixels containing NSB signal)
  Determination of Hillas parameters, parameterising the image ellipse with a width and length parameter, then scaled per telescope and, in stereoscopic systems, the mean over all triggered telescopes is found
  γ-hadron separation: compare the extracted width and length of the shower image of telescope i to the mean values expected from γ -ray MC simulations, given their image amplitude, impact distance, zenith angle, et cetera
Shower image simulation on CHEC (sim_tel_array)

## Slide 30: γ-hadron separation
Shower development simulation for γ- and proton-generated showers
Cherenkov light measure for γ- and proton-generated showers.
Accumulation ring (120 m approx.) for the gamma-induced one
Muon-induced rings in the proton-induced one

## Slide 31: Silicon PhotoMultipliers
Semiconductors consisting of an array of avalanche photodiodes (APDs)
One pixel contains several thousands of diodes (microcells) with a size of few tens of μm
Each microcell: a PN junction supplied with a reverse bias voltage above breakdown (Geiger mode)
A photon or thermal excitation may produce an electron-hole pair in the depleted region
Charge carriers can trigger an avalanche in a microcell  high amplification of the incoming photon signal  one avalanche ∝ one electron-hole pair
Signal of one SiPM pixel: sum of the signal of all microcells, read out over their quenching resistor via a common output
SiPM outline

## Slide 32: Silicon PhotoMultipliers
Main technical challenges
  Temperature dependence (DCR and gain)  temperature stabilization is desirable
  Optical Crosstalk
  Sensitivity in the red
Main advantages
  higher peak-PDE
  magnetic-field independence
  use of moderate voltages (70 V compared to 1000 V of MAPMs or PMTs)
  robustness against high-background illuminations
  better angular response
  simple gain determination using SPE spectra from dark counts
  Simpler manufacture and at a lower cost
  active development of SiPMs

## Slide 33: CTA and SST
CTA differential sensitivity (in Crab Nebula units), D. Depaoli

## Slide 34: Night Sky Background
NSB measured at La Palma, Canary Islands, compared to the Cherenkov spectrum of a gamma-ray induced air shower. (D. Depaoli)

## Slide 35: FW calibration procedure
Evaluate the transmission (1 MHz data)
Evaluate the expected number of photons at camera (1 MHz) via
with k evaluated as
Evaluate the expected number of photons at 600 Hz at camera via
With f evaluated as a scale factor between measurements at different rates

## Slide 36: Considerations: correction on power values
The light coming from the source should be uniform along the sphere
Scan the laser beam along the sphere of radius equivalent to the distance from the source
Our measures: scan on a 2D plane (x-y plane)  some correction is needed
Measures of power have to be corrected by this R factor: in 2D this means and factors

## Slide 37: Correction on power values
The correction factor is evaluated as the ratio between the sensor size on the sphere (CD) and the size it should have to see the same amount of light on the plane (AB)
Since CD = 10.3 mm (real sensor size), we have
Then
and
or
Where (x,y) are the sensor position coordinates at each scan step (we assume they are located at the center of the sensor).

## Slide 38: Correction on power values and final beam model procedure
Final procedure
  Perform a first fit on data without any correction (origin of the sphere is unknown a priori)
  One dimensional fit on the distribution of power along x = 0 and y = 0 with intuitively initialized parameters (function: 1D supergaussian)
  Bidimensional fit on power data with initialization parameters obtained from previous 1D fit (function: 2D supergaussian)
  With the newly found values of coordinates, evaluate the correction factors R and correct the data
  Repeat the 2D fit on corrected data
This procedure can be repeated at anytime by the user giving the new sets of data and experimental premises (via configuration file)
Results from fit are saved both via plots and fit parameters
Distribution of correction factor

## Slide 39: Thorlabs diffusers

## Slide 40: NSB beam
Not interested in knowing the NSB beam perfectly… But the model is working  some considerations
  LED should deliver ≈ monochromatic light, with a peaked spectrum (peak-wavelenght λ = 405 nm): assume it as almost gaussian?
  Same kind of circle tophat diffuser  final result should be a peaked function «flattened» by a tophat again
  Data taken with same method
Tried re-using the supergaussian, expecting a different order n
Initial parameters slightly different
(x0, y0) no longer expected to be (0,0)
MIN and MAX filter cases should be the same, except for the amplitude value
 Possibility to eventually use the same model on NSB beam

## Slide 41: NSB beam
Data @ 1 m, MIN filter
Data @ 3 m, MIN filter
Beam uniformity: 8.11 %
Data @ 1 m

## Slide 42: Additional correction on NSB values
Powermeter always measures at λ = 400 nm
LED light peaked at 405 nm
Data must be corrected
  LED emission cannot be varied
  Wavelength settings on the powermeter can  perform a «wavelength» scan over the entire wavelength range of PM (supplying a constant light)
  Obtain the relative quantum efficiency (with reference value 400 nm) VS wavelength of the sensor
  Correction on power measures will be the variation of QE at 405 nm w.r.t. reference value
Power and Relative quantum efficiency VS wavelength

## Slide 43: Relative QE
Measured power is
By definition
Then
The «peak power» (power @ 400 nm) is
Measure several other power values (@ different λ), while the rate of photoelectrons is always the same
Then
Define the relative quantum efficiency as

## Slide 44: From relative QE to relative responsivity
Definition of responsivity
Relative responsivity
To be compared with data from manufacturer

## Slide 45: CHEC-S
J. Zorn

## Slide 46: CHEC-S
CHEC-S
SST Camera
R.White
32 identical camera modules connected to a single Backplane (BP).
Each camera module contains a SiPM tile with 64-pixels each of active size 6 mm x 6 mm.
Each tile is connected to Front-End Electronics (FEE) module that provides SiPM bias control, waveform digitisation and a first-level camera trigger.

## Slide 47: Pixels’ «health»
1σ criterion
«Bad» pixels: 12.5%
2σ criterion
«Bad» pixels: 2.3%
3σ criterion
«Bad» pixels: 1.5%

## Slide 48: CHEC-S – Dynamic range
Extracted charge VS light from source
Scan over a range of photon numbers
Here: pre-HV reconfiguration
The new HV settings should improve the error bars
Charge_SPE and PDE are not known for the current CHEC-S state (performance after many years and a commissioning)
PDE of MAPM (used in CHEC-M) and of Hamamatsu S10943-1071 SiPM (similar to the one used in CHEC-S) as function of the wavelength. (J. Zorn)

## Slide 49: Longterm tests
Idea: repeat iteratively the fundamental steps of CHEC-S to study its stability
Stability in terms of pixels’ «health», uniformity, response, etc. as a function of time
Monitoring humidity and temperature
To be run ideally for months every night
«Longterm performance» of a prototype
Fundamental steps
  Boot-up: turn on and off m times
  Lid control
  Dynamic range
  Quick analysis and tiding up of data
Repeat n times
A posteriori analysis

## Slide 50: DBM Monitoring: Graphana