## Slide 1: MPIK Update
Isabella Sofia
19/12/2022

## Slide 2: Main activities
Laser beam alignment
Robot arm scans on the diffused (and aligned) beam
Filter wheel calibration procedure (using PM on Robot Arm at 3 m distance)
CHEC-S Longterm tests implementation
Cherenkov images (currently)

## Slide 3: Diffused beam scans
Distance (m) | Diffusers | Beam expander (Y/N) | PM aperture (mm) | Aperture after diffuser | Number of measurements | σ/mean (%) | 
 |  |  |  |  |  | R < 0.1 | Inside camera
3 | 20 + 5 | Y | 10.3 | NO | 10 | 0.93 % | 1.68%
2 | 20 + 5 | Y | 10.3 | NO | 10 | 1.39 % | 2.43 %
1 | 20 + 5 | Y | 10.3 | NO | 10 | 2.93% | 4.77%
D = 2.9983 m
D = 2.0084 m
D = 1.0035 m

## Slide 4: Filter Wheel calibration: procedure
Evaluate the attenuation (1 MHz data)
Evaluate the expected number of photons via the factor k
Evaluate the expected number of photons at 600 Hz: calculate the scale factor between measurements at 1 MHz and 600 Hz at the PM Laser. Mean factor = 0.946
Interpolation on the FW positions that yield at least 1 photon
Save interpolation values in a calibration file for the DBM.

## Slide 5: Example plots

## Slide 6: Other considerations
Few measurements with the SiPM next to the PM Cam on the Robot Arm, to check if the PM is still a reliable device at such a distance.
The variation between the two measurement methods is less then 6% in the region of interest
General agreement between methods: the powermeter is indeed a reliable device to perform such calibrations.
Given that the evaluation of from the SiPM data has been done with a number of approximations (PDE, Crosstalk), we will consider the calibration via the PM for further measurements.
This interpolation results are implemented in the DBM Software.

## Slide 7: Trigger delay scan
Slightly different cable set-up  repeat scan on trigger delays
Delay = 315 ns
Single pulse (pixel n. 0)
Different pixels (bad pixels too)

## Slide 8: Longterm tests
Idea:
  Initial bootup
  Pedestal run
  Dynamic range scan
  Quick analysis and tidying up data
  Analysis of monitoring data (variations with temperature, humidity, versus time…)
Flowchart:
  https://drive.google.com/file/d/1ipderqUFhA6Gs1_d45XclkdSIaL9oYHD/view?usp=sharing

## Slide 9: What’s next
Coming back in January
Data analysis
  Beam scans
  Longterm tests (periodic health check, monitoring data, dynamic range)
  Cherenkov images
…